PM12 Hard Disk Setup

Rev 20110126

This procedure assumes that physical replacement of the Hard Disk has taken place.  If assistance is needed for this, please review the PM12 troubleshooting section of the manual, or call for support, 1-800-274-4212.

When the instrument is first started with the replacement disk

the application will signal "out of service" because of missing

calibration data or has unprogrammed amplifier node addresses.

Since node addresses are different for each machine, the hard drives are set up generically and must be completed using this procedure before the instrument can be operated.

This document describes the necessary steps to complete setup of the hard drive

You will need a USB Keyboard to complete this process.

Make sure that the PM12 Dongle is installed into one of the USB ports during this process

Setup Instructions

I  Node Address Settings

- Power up the PM12 system using the keyswitch

Please note, errors will be displayed on the instrument screen until the steps outlined in this document are completed

If the windows ‘found new hardware’ box appears, press ‘cancel’ on the screen and the instrument will continue to boot

When the system fully boots, ‘Out of Service’ will appear on the screen

- From the PM12 Menu, Access the Administration Mode by pressing on the 'PM12' icon in the bottom left-hand corner of the touch screen.

When the ‘ERROR’ box appears, choose ‘OK’

3. The Windows embedded login screen will appear.

Login with the following credentials:

User name: administrator

Password is: 0mr3hT   caution ‘0’=zero, not the letter O

Note: if additional ‘found new hardware’ messages appear, select ‘cancel’

- From the windows desktop, double tap on the ‘DeviceConfig’ application

From the x channel configuration screen, we will enter the amplifier addresses.

The amplifier addresses will be on a label for each of the two system amplifiers   visible by removing the detector panels. They are a 4 digit number, enter exactly as they appear

Make note of the amplifier addresses:

Amp1____   Amp2____    Amp3____ Amp4____

- Enter the amplifier addresses into the configuration screen as follows:

  - Tap on the Device, Amp 1-4 or BattCharger that you want to change

  - When Highlighted, tap ‘Change Address’

  - Enter the new address number xxxx into the ‘New Address’ line

  - Tap ‘Change Address’

  - Repeat a-d for each amplifier.

  - If the BattCharger is not automatically set to 70, set it at this time following the same steps as done for the amplifiers.

  - After all amplifiers and BattCharger set, tap ‘Save to Config file’

Confirm the save tapping ‘OK’

Note: depending on machine options, the barrier and Lamp box may need to be checked/unchecked and the configuration saved as well.

- After the configuration is saved, Tap ‘exit’ and shut down the system, following \

Standard windows shutdown protocol.

- Power down the PM12 when it is safe using the keyswitch, and then restart after a

few seconds.

8 The machine should power up normally, but may require setup of alarms and calibration Proceed to the next section

II. Alarm Set, Active Calibration and other parameters

Note: There may be default values that allow the system to boot properly if

no errors are noted on boot up. If this is the case, refer to your specific procedures    to confirm instrument settings

- From the PM12 Menu, Access the Admin, ‘Health Physicist’ user Mode by  pressing on the 'PM12' icon in the bottom left-hand corner of the touch screen.

Select the desired user type (Health Physicist) and tap ok.

The Password for Health Physicist is ‘HP’

You may enter the password on the screen keyboard, then tap ‘enter’

- From the setup screen, tap ‘setup’, then ‘alarms’

- Enter values on the ‘Normal’ and ‘Cobalt-60’ lines