WRM2 Serial Data Protocol

|  | Prepared by: | Reviewed by: | Approved by: |
| --- | --- | --- | --- |
| Name | James Lohr | Jeff Bierlink | Dave Jarrow |
| Date |  |  |  |
| Signature |  |  |  |

abcdef

Suite 150

5000 Highlands Parkway

Smyrna, GA 30082

Page 1 of 16

WRM2 Serial Data Protocol

Protocol Description — Revision 3

MGP Instruments, Inc. ©2012

All rights reserved. No part of this document may be reproduced or transmitted in any form or by any means, electronic or mechanical, including photocopying, recording or by any information storage or retrieval system, without permission in writing from MGP Instruments, Inc.

At the time of this release, all information contained in this document is believed to be true and accurate. Any and all comments or corrections received by MGP Instruments, Inc. become the sole property of MGP Instruments, Inc. MGP Instruments, Inc. reserves the right to retain the sole ownership of and to publish any comments and corrections received without any obligation to the sender.

Comments and/or corrections may be sent to:

James Lohr

MGP Instruments, Inc.

Suite 150

5000 Highlands Parkway

Smyrna, Georgia 30082

- Introduction

This document describes the data exchange protocol used between WRM2 Base radios and the host PC and in some cases between dosimeters or other portable instruments and the host using a serial connection.

- References

| Reference # | Description |
| --- | --- |
| 45504-W | LPU/LPDU Parameter Table |
| 45866-BA | RMS Protocol Technical Specifications |

- WRM-2 DOSIMETER SERIAL DATA PROTOCOL

All the data is sent is ASCII format consisting of 46 characters sent at 4800 BAUD from the dosimeter. The baud rate from the WRM2 base is programmable. The other port settings are: no parity, 8 data bits and 1 stop bit. The data starts with a Line Feed (LF) and ends with a checksum CKS.  A Carriage Return (CR) is transmitted after the checksum.

The data sent contains:

Instrument serial number

Accumulated Dose

Current Dose Rate

Instrument Status

Battery Level

For checksum purposes each pair of bytes, not including the <LF> and <CR>, is taken as a Hexadecimal Byte value and considered to be one Byte for the calculation. The check sum is calculated such that the modulo 256 sum of the resultant 23 bytes is equal to 0.

  - 4800 BAUD DMC-2000 & SOR family DOSIMETER Data

| 1 | LF | Line Feed |
| --- | --- | --- |
| 2 | N5 | Dosimeter Serial Number N5N4N3N2N1N0 |
| 3 | N4 |  |
| 4 | N3 |  |
| 5 | N2 |  |
| 6 | N1 |  |
| 7 | N0 |  |
| 8 | Status 1 MSB | Dosimeter’s dose and dose rate alarms and status. |
| 9 | Status 1 LSB |  |
| 10 | D5 | Accumulated Dose D5D4D3D2D1D0.D-1 mrem |
| 11 | D4 |  |
| 12 | D3 |  |
| 13 | D2 |  |
| 14 | D1 |  |
| 15 | D0 |  |
| 16 | D-1 |  |
| 17 | 0 | Transmitter Battery Level (see table) |
| 18 | STATUS 5 (MSB) | Configuration Parameters |
| 19 | STATUS 5 (LSB) |  |
| 20 | X1 | Gamma Dose rate X,YZ * 10T mrem/h |
| 21 | Y1 |  |
| 22 | Z1 |  |
| 23 | T1 |  |
| 24 | Status 3 MSB | Dosimeter dose and dose rate warning alarms and status |
| 25 | Status 3 LSB |  |
| 26 | READER # (MSB) | Last reader seen in marker mode |
| 27 | READER # (LSB) |  |
| 28 | LIM 0 | Limit for Zone 0 |
| 29 | LIM 1 | Limit for zone 1 |
| 30 | STATUS 10 (MSB) | Dosimeter configuration for secondary measurement |
| 31 | STATUS 10 (LSB) |  |
| 32 | COUNTS (MSB) |  |
| 33 | COUNTS (LSB) |  |
| 34 | D5 | Accumulated beta dose if enabled. D5D4D3D2D1D0.D-1 mrem |
| 35 | D4 |  |
| 36 | D3 |  |
| 37 | D2 |  |
| 38 | D1 |  |
| 39 | D0 |  |
| 40 | D-1 |  |
| 41 | 0 | Not used |
| 42 | X2 | Beta Dose rate X,YZ * 10T mrem/h |
| 43 | Y2 |  |
| 44 | Z2 |  |
| 45 | T2 |  |
| 46 | CKS MSB | Checksum |
| 47 | CKS LSB |  |
| 48 | CR | Carriage Return |

  - Transmitter Battery Level

| Byte Value | Meaning |
| --- | --- |
| 0 | Not tested. |
| 1 | 0.9V |
| 2 | 1.0V |
| 3 | 1.1V |
| 4 | 1.2V |
| 5 | 1.3V |
| 6 | 1.4V |
| 7 | 1.5V |
| Any other value | Error |

  - Status Bytes

Status 1: Events

| b7 |  |  | b6 |  | b5 |  | b4 |  | b3 |  | b2 |  | b1 |  | b0 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Battery status 0 =OK 1 = low |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Mode 0 = sat 1 = autonomous |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Rate Saturation 0 = no 1 =yes |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Dose Saturation 0 = no 1 = yes |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Dose Alarm 0 = no 1 = yes |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Rate Alarm 0 = no 1 = yes |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Bat Changed 0 = no 1 = yes |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | General fault 0 = no 1 = yes |  |

Status 3: Alarms

| b7 |  |  | b6 |  | b5 |  | b4 |  | b3 |  | b2 |  | b1 |  | b0 |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Dose warning 1 = exceeded alarm level |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Rate warning 1 = exceeded alarm level |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Time Alarm 1 = exceeded stay time allowed |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Historic saturated 1 = yes |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Integration time saturated 1 = yes (t>99 hours) |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Satellite mode exited 1 = yes |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Internal calib. parameters incorrect 0 = no 1 = yes |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | External calib. parameters incorrect 0 = no 1 = yes |  |  |

Status 5: Operational configuration parameters

| b7 |  |  | b6 | b5 |  | b4 |  | b3 |  | b2 |  | b1 |  | b0 |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Teletransmission mode 0 = without 1 = with |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | flash detection 0 = no 1 = yes |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Teletransmission triggered by 1 = external strobe 0 = periodic ~every 4 sec. |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Teletransmission data speed 0 = 300 BAUD 1 = 4800 BAUD |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Report optical test 0 = no 1 = yes (read only) |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Report Ext. Asic fault :0 = no 1 = yes (read only) |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Report historical fault 0 = no 1 = yes (read only) |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Report E2PROM fault 0 = no 1 = yes (read only) |  |  |

Highlight = default settings

Status 10:

| b7 |  |  | b6 | b5 |  | b4 |  | b3 |  | b2 |  |  | b1 |  | b0 |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Secondary Measurement 0 = without 1 = with |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Display 00 = main measurement 01 = secondary measurement |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 10 = main + secondary 11 = not used |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Main Measurement 0= internal 1= external |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Main Measurement 0= Hp(10) 1= Hp(0.07) |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Secondary Measurement 0= internal 1= external |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Secondary Measurement 0= Hp(10) 1= Hp(0.07) |  |  |
|  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | Validation Control 0= without 1= with |  |  |

Highlight = default settings

- WRM-2 AMP & DRM SERIAL DATA PROTOCOL

All the data is sent in ASCII format consisting of 26 characters sent at 300 BAUD from the AMP or DRM. The baud rate from the WRM2 base is programmable. The other port settings are: no parity, 8 data bits and 1 stop bit. The data starts with a Line Feed (LF) and ends with a checksum CKS.  A Carriage Return (CR) is transmitted after the checksum.

The data sent is:

Instrument serial number

Accumulated Dose (if enabled in some cases)

Current Dose Rate

Instrument Status

For checksum purposes each pair of bytes, not including the <LF> and <CR>, is taken as a Hexadecimal Byte value and considered to be one Byte for the calculation. The check sum is calculated such that the modulo 256 sum of the resultant 12 bytes is equal to 0.

  - AMP & DRM family data

| 1 | LF | Line Feed |
| --- | --- | --- |
| 2 | N5 | Meter ID and Serial Number set via AMPView N5N4N3N2N1N0 |
| 3 | N4 |  |
| 4 | N3 |  |
| 5 | N2 |  |
| 6 | N1 |  |
| 7 | N0 |  |
| 8 | Status | See table below |
| 9 | Extended Status | Always 30 H for AMP, DRM - See table below |
| 10 | D5 | Accumulated Dose or 888888 if dose is disabled D5D4D3D2D1D0.D-1 mrem |
| 11 | D4 |  |
| 12 | D3 |  |
| 13 | D2 |  |
| 14 | D1 |  |
| 15 | D0 |  |
| 16 | D-1 |  |
| 17 |  | 30 H |
| 18 | X | Dose rate X,YZW * 10T mrem/h |
| 19 | Y |  |
| 20 | Z |  |
| 21 | T |  |
| 22 | W |  |
| 23 | 0 | reserved |
| 24 | CKS MSB | Checksum, ASCII translation of HEX CKS |
| 25 | CKS LSB |  |
| 26 | CR | Carriage Return |

Byte 23 – not currently used.

Extended Status

| Value | Meaning |
| --- | --- |
| 30 Hex, 0 ASCII | OK |
| 31 Hex, 1 ASCII | Saturation |
| 32 Hex, 2 ASCII | Rate Alarm |
| 33 Hex, 3 ASCII | External Power Fail |
| 34 Hex, 4 ASCII | Low battery |
| 35 Hex, 5 ASCII | Detector fault |
| 36 Hex, 6 ASCII | Saturation & External Power Fail |
| 37 Hex, 7 ASCII | Saturation & Low battery |
| 38 Hex, 8 ASCII | Saturation & External Power Fail & Low battery |
| 39 Hex, 9 ASCII | Rate Alarm & External Power Fail |
| 41 Hex, A ASCII | Rate Alarm & Low battery |
| 42 Hex, B ASCII | Rate Alarm & External Power Fail & Low battery |
| 43 Hex, C ASCII | External Power Fail & Low battery |
| 44 Hex, D ASCII | External Power Fail & Detector fault |
| 45 Hex, E ASCII | External Power Fail & Low battery & Detector fault |
| 46 Hex, F ASCII | Low battery & Detector fault |
| Any other value | Undefined - Internal error |

Status

| Value | Meaning |
| --- | --- |
| 30 Hex, 0 ASCII | OK |
| 31 Hex, 1 ASCII | Saturation |
| 32 Hex, 2 ASCII | Rate Alarm |
| 33 Hex, 3 ASCII | External Power Fail |
| 34 Hex, 4 ASCII | Low battery |
| 38 Hex, 8 ASCII | Detector fault |
| Any other value | Internal error |

- WRM-2 GENERAL PURPOSE SERIAL DATA PROTOCOL

  - Introduction

A general purpose protocol has been developed for use with the WRM2 radio only. It has no equivalent for direct serial communications. It is used with a ‘smart’ WRM2 transmitter; one that has an embedded reprogrammable microprocessor, for example the RAMSYS WRM2 TRX.

  - The General Packet Wrapper

This wrapper packet can be used to contain any type of ASCII data. The packet is defined as follows:

| 1 | LF | Line Feed |
| --- | --- | --- |
| 2 | A | Packet Type AB = ASCII encode Hex, 00 through FF |
| 3 | B |  |
| 4 | Data | Device data |
| … | Data |  |
| N-5 | Data |  |
| N-4 | W | Checksum WXYZ = compliment of ((Sum of bytes 1 through N-5) Modulo 65536) expressed as 4 Hexadecimal digits |
| N-3 | X |  |
| N-2 | Y |  |
| N-1 | Y |  |
| N | CR | Carriage Return |

Where:

N = 	Length of packet in Bytes

Packet Type:

| Code Number | Packet Type |
| --- | --- |
| 00 | Error |
| 01 | Modbus |
| 02 | RDS Dual Detector |
| 03 | Reserved for future use |
| … |  |
| FF |  |

    - The Modbus Packet

This packet is the data subset of the wrapper packet. This is the protocol used by the RAMSYS WRM2 TRX. It contains a response from the device starting at address 5100 that the transmitter sends verbatim. The balance of the message is filled from other locations read by the transmitter.

For clarity the packet wrapper is included and is grayed out. The Modbus memory that is sent verbatim is shaded in yellow and is described in MGP Document numbers 45866BA and 45504-W.

| 1 | LF | Line Feed |
| --- | --- | --- |
| 2 | 0 | Modbus Packet Type |
| 3 | 1 |  |
| 4 | I | Device ID in requested Hex (01 - FF) |
| 5 | D |  |
| 6 | A | Size of packet: AB Hex or 171 Decimal |
| 7 | B |  |
| 8 | 5 | Modbus Starting Address |
| 9 | 1 |  |
| 10 | 0 |  |
| 11 | 0 |  |
| 12 | I | Device ID sent by device (must match Bytes 4 & 5) See document 45866-BA Section 3.1 |
| 13 | D |  |
| 14 | 0 | 03 Hex: Modbus Read Response See document 45866-BA Section 3.1 |
| 15 | 3 |  |
| 16 | 4 | 44 Hex: Modbus Memory Read Size See document 45866-BA Section 3.1 |
| 17 | 4 |  |
| … |  | Address 5100 Data: see Document 45504-W, section 5.x |
| 154 | Q | Modbus internal Checksum STQR See document 45866-BA Section 2.2 |
| 155 | R |  |
| 156 | S |  |
| 157 | T |  |
| 158 | M | MN = Monitor Type (see section 2.1, document 45504-W) |
| 159 | N |  |
| 160 | a | Algorithm Channel A (see section 2.3, document 45504-W) |
| 161 | a |  |
| 162 | b | Algorithm Channel B (see section 2.3, document 45504-W) |
| 163 | b |  |
| 164 | c | Algorithm Channel C (see section 2.3, document 45504-W) |
| 165 | c |  |
| 166 | d | Algorithm Channel D (see section 2.3, document 45504-W) |
| 167 | d |  |
| 168 | W | Checksum WXYZ = compliment of ((Sum of bytes 1 through N-5) Modulo 65536) |
| 169 | X |  |
| 170 | Y |  |
| 171 | Y |  |
| 172 | CR | Carriage Return |

    - The RDS Dual Detector Packet

This packet is the data subset of the wrapper packet. This is the protocol used by the RDS when firmware enabled to transmit the second detector information. It uses the WRM2 AMP and DRM Protocol when setup to send only one detector.

For clarity the packet wrapper is included and is grayed out. The RDS data is shown in white.

| 1 | LF | Line Feed |
| --- | --- | --- |
| 2 | 0 | Packet Type |
| 3 | 2 |  |
| 4 | N5 | RDS Serial Number N5N4N3N2N1N0 |
| 5 | N4 |  |
| 6 | N3 |  |
| 7 | N2 |  |
| 8 | N1 |  |
| 9 | N0 |  |
| 10 | Status 1 | See Table Status 1 below |
| 11 | Extended Status 1 | See Table Extended Status 1 below |
| 12 | D5 | Accumulated Dose Detector 1 D5D4D3D2D1D0.D-1 |
| 13 | D4 |  |
| 14 | D3 |  |
| 15 | D2 |  |
| 16 | D1 |  |
| 17 | D0 |  |
| 18 | D-1 |  |
| 19 | X1 | Gamma Dose rate Detector 1 X,YZW* 10T |
| 20 | Y1 |  |
| 21 | Z1 |  |
| 22 | T1 |  |
| 23 | W1 |  |
| 24 | M1 | Units of Measure Detector 1 (See Table Below) |
| 25 | D5 | Accumulated dose Detector 2 D5D4D3D2D1D0.D-1 |
| 26 | D4 |  |
| 27 | D3 |  |
| 28 | D2 |  |
| 29 | D1 |  |
| 30 | D0 |  |
| 31 | D-1 |  |
| 32 | X2 | Dose Rate Detector 2 X,YZW * 10T |
| 33 | Y2 |  |
| 34 | Z2 |  |
| 35 | T2 |  |
| 36 | W2 |  |
| 37 | M2 | Units of Measure Detector 2 (See Table Below) |
| 38 | Status 2 | See Table Status 2 below |
| 39 | W | Checksum WXYZ = compliment of ((Sum of bytes 1 through N-5) Modulo 65536) |
| 40 | X |  |
| 41 | Y |  |
| 42 | Y |  |
| 43 | CR | Carriage Return |

The Units of Measure (valid values 0-9 and A-F)

| Value for M1 and M2 | Dose Units | Dose Rate Units |
| --- | --- | --- |
| 0 (0x30) | M1 = error, M2 = no detector | M1 = error, M2 = no detector |
| 1 (0x31) | mrem | mrem/h |
| 2 (0x32) | rem | rem/h |
| 3 (0x33) | mSv | mSv/h |
| 4 (0x34) | Sv | Sv/h |
| 5 (0x35) | N/A | Counts Per Minute |
| 6 (0x36) | N/A | DPM |
| 7 (0x37) | N/A | Counts Per Second |
| 8 (0x38) | N/A | Bq |
| 9 (0x39) | N/A | Bq/cm2 |
| A (0x41) | µSv | µSv/h |
| B (0x42) | µrem | µrem/h |
| C through F | reserved / error | reserved / error |
| Any other value | error | error |

Status 1

| Value | Meaning |
| --- | --- |
| 0 (0x30) | OK |
| 1 (0x31) | Overflow |
| 2 (0x32) | Dose Rate Alarm |
| 3 (0x33) | No External Power |
| 4 (0x34) | Low battery |
| 5 (0x35) | Detector Fault |
| 6 (0x36) | Internal Error (3.3V or HV) |
| 7 (0x37) | Program CRC |

Status 2

| Value | Meaning |
| --- | --- |
| 0 (0x30) | OK |
| 1 (0x31) | Overflow |
| 2 (0x32) | Dose Rate Alarm |
| 5 (0x38) | Detector Fault |

Extended Status 1

| Value | Meaning |
| --- | --- |
| 0 (0x30) | OK |
| 1 (0x31) | Saturation |
| 2 (0x32) | Dose Rate Alarm |
| 3 (0x33) | No External Power |
| 4 (0x34) | Low battery |
| 5 (0x35) | Detector Fault |
| 6 (0x36) | Saturation and External Power Fail |
| 7 (0x37) | Saturation and Low Battery |
| 8 (0x38) | Saturation and External Power Fail and Low Battery |
| 9 (0x39) | Rate Alarm and No External Power |
| A (0x41) | Rate Alarm and low Battery |
| B (0x42) | Rate Alarm and No External Power and Low battery |
| C (0x43) | No External Power and Low battery |
| D (0x44) | No External Power and Detector fault |
| E (0x45) | No External Power and Low battery and Detector fault |
| F (0x46) | Low Battery and Detector Fault |
| G (0x47) | Internal Error |

- CRC Calculation

  - Introduction

Except for the WRM2 General Purpose Packet, the CRC comes from the device itself. The data from the device is converted to ASCII by the transmitter. Therefore, to calculate the checksum, we must convert the data back into Hexadecimal before summing the bytes.

For the WRM2 General Purpose Packet, the same convention is used, except the checksum is 4 Bytes instead of 2.

  - Sample Code

    - All Packet Types except WRM2 General Purpose Packet

The following Basic code checks the Checksum for validity and returns true or false for a good checksum of bad checksum respectively.

Private Function CheckCRC(Inbuff() As Byte, Length As Integer) As Boolean

Dim i As Long

Dim crc As Long

Dim CS As Long

CheckCRC = False

For i = 1 To Length Step 2

crc = (crc + Val("&h" & Chr(Inbuff(i)) & Chr(Inbuff(i + 1)))) Mod 256

Next i

If crc = 0 Then CheckCRC = True

End Function

    - WRM2 General Purpose Packet

The following Basic code checks the Checksum for validity and returns true or false for a good checksum of bad checksum respectively.

Private Function CheckModbusWRM2CRC(Inbuff() As Byte, Length As Integer) As Boolean

Dim sum As Long

Dim i As Integer

Dim s As String

Dim l As Long

sum = 0

CheckModbusWRM2CRC = False

For i = 0 To Length - 5

sum = sum + Inbuff(i)

Next i

sum = sum Mod 65536

s = "&H" & Chr(Inbuff(Length - 4)) & Chr(Inbuff(Length - 3)) & Chr(Inbuff(Length - 2)) & Chr(Inbuff(Length - 1))

l = CLng(s)

If l + sum = 65535 Then CheckModbusWRM2CRC = True

End Function